## Slide 1: Method Modifications & ATP 40 CFR part 136.6 and 136.5
OELA Meeting 5/24/2017
Scott Hoatson
DEQ Agency Quality Assurance Officer
DEQLabTechSupport@deq.state.or.us

## Slide 2: 40 CFR 136.6 – Method Modifications
Some methods are old, technology changes, regulations change
Matrices are complicated
Need lower reporting limits
What are your options?

## Slide 3: 40 CFR 136.6 – When is it OK?
Acceptable reasons for an analyst to modify a method include analytical practices
  that lower detection limits
  improve precision
  reduce interferences
  lower laboratory costs
  promote environmental stewardship by reducing generation of laboratory wastes.
Notes: Note: SM generally does not allow modifications to their methods.

## Slide 4: Method Modifications - Good
Maintain Chemistry of the method
Same Determinative Technique
Achieve equivalent performance

## Slide 5: Acceptable Method Modifications
Changes between manual method, flow analyzer, and discrete instrumentation (same basic chemistry)
Changes between automated and manual sample preparation, such as digestions, distillations, and extractions; in-line sample preparation is an acceptable form of automated sample preparation for CWA methods
Interference reductions for ICPMS (e.g collision cell technology)

## Slide 6: More Acceptable Method Modifications
  The use of EPA Method 200.2 or the sample preparation steps from EPA Method 1638, including the use of closed-vessel digestion, is allowed for EPA Method 200.8,
  Changes in pH adjustment reagents. Changes in compounds used to adjust pH are acceptable as long as they do not produce interference.
  Prepackaged reagents

## Slide 7: More Acceptable Method Modifications
  Changes in buffer reagents are acceptable provided that the changes do not produce interferences.
  Changes in the order of reagent addition are acceptable provided that the change does not alter the chemistry and does not produce an interference
  Changes in calibration range (provided that the modified range covers any relevant regulatory limit and the method performance specifications for calibration are met)

## Slide 8: More Acceptable Method Modifications
  Changes in calibration model. (A) Linear calibration models do not adequately fit calibration data with one or two inflection points.
  At least 6 points for a Quadratic Curve
  Cannot extrapolate past highest point.
  Can use weighted curve fits
  The RSE may be used as an alternative to correlation coefficients and coefficients of determination for evaluating calibration curves for any of the methods at Part 136.
    RSE is based on difference of each cal point from true value.
Notes: Example Anions by IC 300.0

## Slide 9: More Acceptable Method Modifications
  Changes in equipment such as equipment from a vendor different from the one specified in the method.
  The use of micro or midi distillation apparatus in place of macro distillation apparatus.
  The use of digital titrators and methods where the underlying chemistry used for the determination is similar to that used in the approved method.

## Slide 10: More Acceptable Method Modifications
  Use of selected ion monitoring (SIM) mode for analytes that cannot be effectively analyzed in full-scan mode and reach the required sensitivity.
  Change purge and trap volumes to lower detection levels
  Change purge gas, Increase temperature during purge, Trap sorbent, desorb time
Notes: There are caveats need to read.

## Slide 11: The Pumpkin Book
Solutions to Analytical Chemistry Problems with Clean Water Act Methods
  EPA 821-R-07-002 March 2007

## Slide 12: Pumpkin Book - Continued
Sampling Requirements
Flexibility to Modify Methods
Data Required to Document Matrix Interference
Solutions to Matrix Interference Problems
When a Matrix Interference is Demonstrated
And much more

## Slide 13: Minimize the Effects
Extraction emulsions
  Use Liquid-Liquid extraction
  Solid Phase extraction
Metals
  Sample Digestion even if Turbidity is low
  Collision Cell ICP-MS
  Use of MSA
  Improve house keeping to reduce contamination
  Use IC for Chromium VI (218.6)

## Slide 14: Minimize the Effects
Organics
  Alternative method
  Isotope dilution
  pH adjustment - separate extractions for acids and bases
  High Volume Injectors are OK
    High vol injector may allow you to dilute and still meet QLs

## Slide 15: Minimize the Effects
  Clean-ups
    GPC – removes lipids and high molecular weight interferences – 1625, 625
    Solid Phase Extraction – Interferences Pesticides and 625
    Florisil, Alumina, and Silica Gel – Separate neutral species from polar interferences.
  Isotope dilution – EPA 1625

## Slide 16: All for one and one for all?
Matrix issues are often related to specific analytes.
If dilution IS required for specific analytes, report as much as possible from the undiluted analysis and ONLY those problem analytes from the dilution.

## Slide 17: Plan B - Alternate Test Procedure
40 CFR 136.5 Alternate Test Procedure
  Limited use
  Facility specific

## Slide 18: EPA Region 10 – Limited Use ATP Protocols
“Protocol for Review and Validation of Alternate Test Procedures for Regulated Organic and Inorganic Analytes in Wastewater Under EPA’s Alternate Test Procedure Program.” Feb 2016 https://www.epa.gov/sites/production/files/2016-03/documents/chemical-atp-protocol_feb-2016.pdf

## Slide 19: EPA Region 10 – Limited Use ATP Protocols
“EPA Microbiological Alternate Test Procedure (ATP) Protocol for Drinking Water, Ambient Water, Wastewater, and Sewage Sludge Monitoring Methods,” September 2010
https://www.epa.gov/sites/production/files/2015-09/documents/micro_atp_protocol_sept-2010.pdf

## Slide 20: EPA Region 10 - ATP
  First, check 40 CFR 136.6 approved method modifications
Working on SOP to make process more clear
  Better description of what needs to be submitted and how.

## Slide 21: EPA Region 10 – ATP Tips
Send to Donald M. Brown R10 QA Manager
  Brown.DonaldM@epa.gov (206) 553-0717
  1200 6th Ave, OERA-140, Seattle, WA 98101
  Operating under 2016 MUR
    Not officially promulgated but approved by previous Director.
    ATP and Approved Methods

## Slide 22: EPA Region 10 – ATP Tips
  One Modification per ATP Application
  Avoid bundling several modifications (especially unrelated ones)
  Include official ATP forms
    Have application package well organized and well labeled
    Raw data clearly associated with any summaries
  Allow plenty of lead time and have patience – Region 10 is currently resource challenged.

## Slide 23: Questions

## Slide 24: EPA FEM Webinar
“Flexible Approaches to Environmental Measurements”
Goals
  Increased emphasis on flexibility in choosing sampling and analytical approaches to meet regulatory requirements for measurements

## Slide 25: Goals
  Development of processes for validation that confirm that measurements meet quality requirements
  Increased collaboration with stakeholders to develop validation processes for new measurement technology
  Rapid assessment of new or modified technologies, methods and procedures

## Slide 26: More content
How these flexible approaches can be applied to different programs, such as air monitoring, CWA and SDWA (e.g., ATP process), RCRA methods, pesticide analysis, etc.
June 20th 8:00am - 9:30am PDT
July 10th 10:00am – 11:30am PDT

## Slide 27: Webinar Information
The teleconference line:
  866-299-3188/5874952464# and
  Adobe Connect link:
  http://epawebconferencing.acms.com/fem-flex2017/